Introduction

In this project, we are going to design a circuit for measuring the temperature using the LM324 op-amp, thermistor, AVR ATmega328 microcontroller & display the reading on the 16×2 LCD display as temperature. After reading the temperature, LM324 op-amp gives the analog output to the AVR ATmega328 microcontroller.

The analog output from LM324 op-amp is sent to the ADC of the ATmega328 that converts the analog data to the digital by inbuilt ADC in it.

This project is very useful for controlling of temperature, It can be used for soldering station to make it automatic temperature controlled & as much at homes as in industrial applications.

Circuit Operation & Design Procedure

A basic low-parts-count bridge circuit is used for this project. Resistors R1 are 1% low Tempco (metal-film) resistors. Current I is constant & set by R1, Rref & E. That is, I= E/(R1+Rref). Note that thermistor current is constant & equal to I because the voltage drops across both R1 & R2 resistors are equal (Ed=0V).

- To select any thermistor on a trial basis. The thermistor converts a temperature change to a resistance change.

- To select the reference temperature. At the reference temperature, V0 must equal zero. To select either the low limit of 200C  or the high limit of 800C. We selected the low limit of 200C for this project. We have just defined Rref. Rref is equal to thermistors resistance at the reference temperature. Specifically, Tref=250C; therefore, Rref=16.6 kΩ. Now calculate ΔR for each temperature.

- To predict the voltage-temperature characteristics. We selected the bridge circuit because it converts a resistance change ΔR into an output voltage.

- To select resistors R1, R2to equal 10 kΩ.

- To make a trial for choice for E=5.0V.

- Calculate I from I=E/(Rref+R1) = 5V/(16.6 kΩ+10 kΩ) = 0.188mA

- Calculate V0 for each value of R from equation V0=IΔR.

- To document performance. V0 is plotted against temperature.

Figure No.-1: An input temperature change of 200C to 800C gives an output voltage       of 0 to 2.914V.

Procedure of Calibration (with data table)

The resistor from (+) input to ground is always chosen to equal the reference resistance of the thermistor. We wanted to V0 to be zero volts when R20=Rref. This will allow us to calibrate or check operation of the bridge.

Calculation for Temperature-to-Voltage Converter is given below:

| Temp(0C) | Rtherm (kΩ) | ΔR(Ω) | V0(volt) |
| --- | --- | --- | --- |
| 20 | 16.6 | 0 | 0 |
| 25 | 12.9 | 3.70 | 0.6956 |
| 30 | 10.3 | 6.30 | 1.1844 |
| 35 | 8.10 | 8.50 | 1.5980 |
| 40 | 6.10 | 10.50 | 1.9740 |
| 45 | 4.80 | 11.80 | 2.2184 |
| 50 | 3.90 | 12.70 | 2.3876 |
| 55 | 3.10 | 13.50 | 2.5380 |
| 60 | 2.50 | 14.10 | 2.6508 |
| 65 | 2.10 | 14.50 | 2.7260 |
| 70 | 1.65 | 14.95 | 2.8106 |
| 75 | 1.30 | 15.30 | 2.8764 |
| 80 | 1.10 | 15.50 | 2.9140 |

Component Description

LM7805: Voltage regulator IC maintains the output voltage at a constant value. &805 IC, a voltage regulator integrated circuit (IC) is a member of 78xx series of fixed linear voltage regulator ICs used to maintain such fluctuations. The xx in 78xx indicates the fixed output voltage it provides.

Potentiometer (50kΩ): A potentiometer is a manually, adjustable, variable resistor with three terminals. Two terminals are connected to a resistive element, the third terminal is connected to an adjustable wiper. The position of the wiper determines the output voltage.

Thermistor (10kΩ): A thermistor is a resistance thermometer, or a resistor whose resistance is dependent on temperature. The term is a combination of “thermal” & “resistor”. It is made of metallic oxides, pressed into a bead, disk or cylindrical shape & then encapsulated with an impermeable material such as epoxy or glass.

LM324: LM324 is a 14pin IC consisting of four independent operational amplifiers (op-amps)   compensated in a single package. Op-amps are high gain electronic voltage amplifier with differential input and, usually, a single-ended output. They can be used as amplifiers, comparators, oscillators, rectifiers etc.

Crystal Oscillator: A crystal oscillator is an electronic oscillator circuit that uses the mechanical resonance of a vibrating crystal of piezoelectric material to create the electrical signal with a precise frequency. Quartz crystals are manufactured for frequencies from a few tens of kilohertz to hundreds of megahertz.

ATmega328: The ATmega328 is a very popular microcontroller chip produced by Atmel. It is an 8-bit microcontroller that has 32k of flash memory, 1k of EEPROM, and 2k of internal SRAM. ATmega328 has 28 pins. The device operates between 1.8-5.5 volts. It provides a highly flexible & cost effective solution to many embedded control applications.

LCD (16×2): 16×2 LCD unit is widely using in embedded system projects because it is cheap, easily available, small in size & easy to interface. 16×2 have two rows & 16 columns, which means it consist 16 blocks of 5×8 dots. 16 pin for connections in which 8 data bits D0-D7 and 3 control bits namely RS, RW & EN. Rest of pins are used for supply, brightness control & for backlight.

Resistors:  We have used three 10 kilo ohm, one 1 kilo ohm & one 100 ohm resistors in this project.

Battery:  Two 9V DC battery source are used to power the circuit.

Total Cost

| Components Name | Price (Taka) |
| --- | --- |
| LM324 | 15 |
| Resistors (5 Pieces) | 5×1=5 |
| ATmega328 | 190 |
| LM7805 | 15 |
| 9V DC Battery (2 Piece) | 25×2=50 |
| Potentiometer | 35 |
| Connecting Wires | 30 |
| Thermistor | 10 |
| Crystal Oscillator | 15 |
| Battery Cap | 15×2=30 |
| LCD | 165 |
| Breadboard | 120×2=240 |
| Report Printing & Binding | 80 |
| Total Cost | 880 |

Conclusion

By this project, we learnt that how temperature is measured using microcontroller & show the reading on LCD display as well as how a bridge circuit converts the resistance change of a thermistor into a voltage change. The circuit output voltage is linear with respect to ΔR.

However, ΔR is not linear with respect to temperature. Therefore, V0 is not linear with respect to temperature. The bridge simply transmits the nonlinearity of the thermistor.

The sensitivity of the temperature-to-voltage converter can be increased easily by increasing E. The maximum value of E is set by the maximum thermistor current to avoid self-heating.